Standard Operating Procedure for Disinfection of COVID 19 Ventilators

(Version 1)

Prepared by

Hospital Infection Control Committee

King George’s Medical University, Lucknow

Aim: To ensure that the staff working at Covid-19 Ventilators, KGMU is adopting appropriate practices to disinfect ventilator units.

Scope: Cleaning and disinfection of COVID-19 ventilator surfaces/units.

General Instructions:

- Wear complete PPE before cleaning and disinfecting the ventilator surfaces and parts

- Parts of ventilator surfaces to be cleaned and disinfected

  - Ventilator monitor

  - Ventilator plug and power cord

  - Dust filters (of fan/ main air outlet/ air intake)

  - Ventilator tubing’s- inspiratory and expiratory (or use disposable)

  - Inspiratory and expiratory valves

  - Humidifier (autoclavable/non-autoclavable)

  - Ambu bag

  - Reusable mask for non-invasive ventilation

  - Other tubing surfaces

  - Laryngoscope

  - Suction device (discard disposable parts)

  - Bain circuit- disposable

  - T piece- disposable

  - Nebuliser bottle

- Doff PPE after every cleaning and disinfection procedure in the doffing area and discard PPE in appropriate bin.

- Discard disposable ventilator parts in appropriate bin

Disinfection Protocols of ventilator surfaces/parts:

| S.No. | Type of surface/ part of ventilator/ suction machine | Cleaning | Disinfection | Frequency |
| --- | --- | --- | --- | --- |
|  | Ventilator monitor/ surfaces | Wipe with detergent water followed by water | Wipe with accelerated H2O2 wipes (BS hydro) | Three times daily |
|  | Ventilator plug and power cord (Do not unplug while being used on patient) | Wipe with detergent water followed by water | Wipe with accelerated H2O2 wipes (BS hydro) | Three times daily |
|  | Dust filters (of fan/ main air outlet/ air intake) | Dip in detergent water followed by soaking in water and air dry | Dip in 1% Bleach for 30 minutes followed by rinsing with water | Every 4 weeks or when necessary |
|  | Ventilator tubing’s- inspiratory and expiratory (or use disposable) | For reusable tubes- wash with detergent and water, both the lumen and outer surfaces followed by rinsing thoroughly in water | Dip in 1% Bleach for 30 minutes followed by rinsing with water | After 5 days |
|  | Inspiratory and expiratory valves | Wipe with detergent water followed by water | Wipe with accelerated H2O2 wipes (BS hydro) | After patient use |
|  | Humidifier (autoclavable/non autoclavable) | Wash using soft scrub or brush with detergent and water followed by rinsing in water | Non-autoclavable: Dip in activated 2.4% Glutaraldehyde for 20 minutes for high level disinfection or for 6 hours for chemical sterilisation Autoclavable: Autoclave as per manufacturer instructions | After every 48 hours |
|  | Ambu bag | Same as humidifier | Same as humidifier | After every patient use |
|  | Reusable mask | Same as humidifier | Same as humidifier | After 48 hours |
|  | Other tubing surfaces | Wipe with detergent and water followed by water and air dry | Wipe with accelerated H2O2 wipes (BS Hydro) | Thrice Daily |
|  | Laryngoscope (except the laryngoscope bulb) Laryngoscope with video (Dispose disposable parts such as blades) | Dip in detergent and water followed by rinsing in water. Clean lumen with brush | Dip in activated 2.4% Glutaraldehyde for 20 minutes for high level disinfection (HLD) or for 6 hours for chemical sterilisation or autoclave Wipe the laryngoscope bulb with 70% alcohol swab | After every patient use |
|  | Suction device (discard disposable parts/ never submerge the electrical parts) | Wipe the electrical parts with detergent and water and air dry Reusable jars must be scrubbed with detergent water and rinsed with water Discard disposable tubing’s | Wipe with H2O2 wipes Reusable jars/ tubing’s should be autoclaved if autoclavable or dipped in activated 2.4% glutaraldehyde for 20 minutes (HLD) or 6 hours for chemical sterilisation | When 1/3rd full |
|  | Nebuliser bottle | Same as humidifier | Same as humidifier | Every 48 hours/ or early if necessary |

References:

- Guidelines on Infection prevention and control of epidemic and pandemic prone acute respiratory infections in health care. World Health Organisation. 2014

- Expert Consensus of Mechanical Ventilation Control for Novel Coronavirus. Chinese Journal of Respiratory and Critical Care Medicine. 2020

- Clinical Management of severe acute respiratory infection when novel coronavirus (2019-nCoV) infection is suspected. Interim Guidance. World Health Organisation. 2020

Check list:

| S.No. | Type of surface/ part of ventilator/ suction machine | Cleaning | Disinfection | Frequency |
| --- | --- | --- | --- | --- |
|  | Ventilator monitor/ surfaces | Yes No | Yes No | Date: |
|  | Ventilator plug and power cord (Do not unplug while being used on patient) | Yes No | Yes No | Date: |
|  | Dust filters (of fan/ main air outlet/ air intake) | Yes No | Yes No | Once a month Date: |
|  | Ventilator tubing’s- inspiratory and expiratory (or use disposable) | Yes No | Yes No | After 5 days Date: |
|  | Inspiratory and expiratory valve surfaces | Yes No | Yes No | After patient use Date: |
|  | Humidifier (autoclavable/non autoclavable) | Yes No | Yes No | After every 48 hours Date: |
|  | Ambu bag | Yes No | Yes No | After every use Date: |
|  | Resusable mask | Yes No | Yes No | After 48 hours |
|  | Other tubing surfaces | Yes No | Yes No | Thrice Daily |
|  | Laryngoscope (except the laryngoscope bulb) Laryngoscope with video (Dispose disposable parts such as blades) | Yes No | Yes No | After every use |
|  | Suction device (discard disposable parts/ never submerge the electrical parts) | Yes No | Yes No | When 1/3rd full Date: |
|  | Nebuliser bottle | Yes No | Yes No | Every 48 hours/ or early if necessary Date: |